## 00_Cover
DIAGNOSTIC CODEBOOK — END TO END
ESP32 + STM32 · bucket-v1 · 4×32 KB flash · HTTPS upload · format_ver=3
How to read this workbook
1. Readable code = PREFIX-NN (decimal). Example: family 0x18 + code 16 → PM-16.
2. Faults (STM 0x59 / ESP diag_report): session+sequence identify the event; no numeric 'value'.
3. Metrics (PM / SD / SSID chunks): flags bit 0x04 set; value_u32 = source_session ¦ (source_sequence << 16).
4. Path: detect → 16-byte flash record → fill 32 KB bucket → seal → HTTPS POST 32768 bytes → HTTP 2xx → erase.
5. STM 0x5B status (no ACK) → ESP stores SD-01…SD-16 → server concatenates → 74-byte field map.
6. WiFi SSID name: PM-29 (len) + PM-30…37 (8×u32 LE) → ASCII string; no NVS map needed on server.
7. Sheets: Families · All_Codes (filterable master) · STM_Faults · ESP_Faults · PM_Metrics · SD_Status · Fields · Flow.
Legend — row colors in code sheets
CRITICAL
ERROR / fault
WARN
METRIC / INFO
SSID chunk
Totals: 25 families · 175 STM faults · 45 ESP faults · 49 PM/SSID · 19 SD · 288 codes overall
Sources: DIAG_SERVER_REGISTRY.md · DIAG_HTTPS_SERVER_API.md · COMPONENT_DIAGNOSTIC_ANALYSIS.md · main/diag_registry.h

## 01_Families
FAMILIES
Shared family IDs — STM owns 0x01–0x0F · ESP owns 0x10–0x19
family_id | hex | PREFIX | Owner | Area / meaning | Record style | UART / source
1 | 0x01 | ST | STM | Framing & command contract | fault / alert | STM CMD 0x59 (alerts)
2 | 0x02 | SU | STM | STM↔ESP UART (USART3) | fault / alert | STM CMD 0x59 (alerts)
3 | 0x03 | SA | STM | STM→ESP audio / SPI transport | fault / alert | STM CMD 0x59 (alerts)
4 | 0x04 | SF | STM | Stream / application flow | fault / alert | STM CMD 0x59 (alerts)
5 | 0x05 | SC | STM | Cloud / WiFi status via ESP | fault / alert | STM CMD 0x59 (alerts)
6 | 0x06 | SS | STM | STM runtime / faults / reset | fault / alert | STM CMD 0x59 (alerts)
7 | 0x07 | SO | STM | OTA control over UART | fault / alert | STM CMD 0x59 (alerts)
8 | 0x08 | PW | STM | MP2722 / battery / charging / NTC | fault / alert | STM CMD 0x59 (alerts)
9 | 0x09 | IB | STM | I2C bus / controller | fault / alert | STM CMD 0x59 (alerts)
10 | 0x0A | TM | STM | Temperature probes | fault / alert | STM CMD 0x59 (alerts)
11 | 0x0B | AH | STM | DAC / SAI / MDF / microphones | fault / alert | STM CMD 0x59 (alerts)
12 | 0x0C | DP | STM | ST7789V2 display | fault / alert | STM CMD 0x59 (alerts)
13 | 0x0D | FM | STM | MX25L6433F external flash | fault / alert | STM CMD 0x59 (alerts)
14 | 0x0E | UB | STM | USB host / SpO2 | fault / alert | STM CMD 0x59 (alerts)
15 | 0x0F | IO | STM | Buttons / GPIO / power rails | fault / alert | STM CMD 0x59 (alerts)
16 | 0x10 | PR | ESP | Provisioning / pairing | fault / alert | ESP local
17 | 0x11 | UA | ESP | UART bridge / STM link | fault / alert | ESP local
18 | 0x12 | SP | ESP | SPI audio path | fault / alert | ESP local
19 | 0x13 | WF | ESP | Wi‑Fi faults | fault / alert | ESP local
20 | 0x14 | TL | ESP | TLS / reachability faults | fault / alert | ESP local
21 | 0x15 | MQ | ESP | MQTT faults | fault / alert | ESP local
22 | 0x16 | ES | ESP | ESP system / crash / heap thresholds | fault / alert | ESP local
23 | 0x17 | OT | ESP | OTA faults | fault / alert | ESP local
24 | 0x18 | PM | ESP | Performance / link metrics (integers) | metric chunks | ESP health / WiFi / MQTT / ping / DNS / TLS
25 | 0x19 | SD | ESP | STM 0x5B status dump chunks | metric chunks | STM CMD 0x5B → ESP SD-01…19

## 02_All_Codes
ALL CODES (MASTER)
Filter by Owner / Kind / Severity / PREFIX · Full end-to-end registry
Code | family_id | hex | PREFIX | code_id | Owner | Kind | Severity | Unit | Meaning (what it means) | When / path | Value decode | Notes
ST-01 | 1 | 0x01 | ST | 1 | STM | fault | ERROR | — | LEN is below 2 or above the protocol maximum; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-02 | 1 | 0x01 | ST | 2 | STM | fault | ERROR | — | END byte is not 0x55; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-03 | 1 | 0x01 | ST | 3 | STM | fault | ERROR | — | received XOR does not match calculated XOR; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-04 | 1 | 0x01 | ST | 4 | STM | fault | CRITICAL | — | raw RX buffer would overflow; clears after parser reset and a valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-05 | 1 | 0x01 | ST | 5 | STM | fault | WARN | — | START arrives during a partial frame and forces resynchronization; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-06 | 1 | 0x01 | ST | 6 | STM | fault | ERROR | — | partial frame exceeds assembly timeout; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-07 | 1 | 0x01 | ST | 7 | STM | fault | WARN | — | valid frame contains unsupported command ID; clears on supported command. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-08 | 1 | 0x01 | ST | 8 | STM | fault | ERROR | — | payload length violates command contract; clears on valid instance of that command. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-09 | 1 | 0x01 | ST | 9 | STM | fault | ERROR | — | nested length exceeds payload bounds; clears on valid command payload. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-10 | 1 | 0x01 | ST | 10 | STM | fault | ERROR | — | STM tries to build an oversized TX payload; clears after valid TX build. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-11 | 1 | 0x01 | ST | 11 | STM | fault | CRITICAL | — | TX checksum self-verification fails; clears after verified build. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-01 | 2 | 0x02 | SU | 1 | STM | fault | ERROR | — | USART3 overrun (ORE); clears after flag recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-02 | 2 | 0x02 | SU | 2 | STM | fault | ERROR | — | framing error (FE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-03 | 2 | 0x02 | SU | 3 | STM | fault | WARN | — | noise error (NE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-04 | 2 | 0x02 | SU | 4 | STM | fault | ERROR | — | parity error (PE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-05 | 2 | 0x02 | SU | 5 | STM | fault | WARN | — | receiver timeout/IDLE with incomplete frame; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-06 | 2 | 0x02 | SU | 6 | STM | fault | ERROR | — | USART3 RX DMA error if DMA is enabled later; clears after recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-07 | 2 | 0x02 | SU | 7 | STM | fault | ERROR | — | USART3 TX DMA error if DMA is enabled later; clears after successful TX. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-08 | 2 | 0x02 | SU | 8 | STM | fault | ERROR | — | finite-timeout packet TX returns HAL error/timeout; clears after successful TX. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-09 | 2 | 0x02 | SU | 9 | STM | fault | ERROR | — | ESP is silent for a generic request expecting a reply; clears when valid ESP traffic resumes. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-10 | 2 | 0x02 | SU | 10 | STM | fault | ERROR | — | validated RX queue is full and packet would be lost; clears after queue recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-01 | 3 | 0x03 | SA | 1 | STM | fault | ERROR | — | SPI2 audio transfer returns non-timeout HAL error; clears after successful block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-02 | 3 | 0x03 | SA | 2 | STM | fault | ERROR | — | SPI2 audio transfer times out; clears after successful block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-03 | 3 | 0x03 | SA | 3 | STM | fault | ERROR | — | SPI2 audio DMA error if enabled; clears after recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-04 | 3 | 0x03 | SA | 4 | STM | fault | WARN | — | PCM transport queue is full and block is dropped; clears after sustained drain. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-05 | 3 | 0x03 | SA | 5 | STM | fault | ERROR | — | PCM block is null, zero, oversized, or misaligned; clears on valid block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-01 | 4 | 0x04 | SF | 1 | STM | fault | ERROR | — | illegal stream state transition; clears after returning to IDLE/STOPPED. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-02 | 4 | 0x04 | SF | 2 | STM | fault | ERROR | — | START response timeout; clears on correlated success or cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-03 | 4 | 0x04 | SF | 3 | STM | fault | ERROR | — | START response reports failure; clears on later success/cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-04 | 4 | 0x04 | SF | 4 | STM | fault | ERROR | — | START response pet/audio/state fields do not match request; clears on correlated response. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-05 | 4 | 0x04 | SF | 5 | STM | fault | ERROR | — | STOP response timeout; clears on correlated success or forced reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-06 | 4 | 0x04 | SF | 6 | STM | fault | ERROR | — | STOP response reports failure; clears on later success/reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-07 | 4 | 0x04 | SF | 7 | STM | fault | ERROR | — | STOP response does not correlate; clears on correlated response. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-08 | 4 | 0x04 | SF | 8 | STM | fault | WARN | — | STM and ESP stream states diverge; clears when both agree. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-09 | 4 | 0x04 | SF | 9 | STM | fault | WARN | — | screen/mode exits while handshake is pending; clears after settlement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-01 | 5 | 0x05 | SC | 1 | STM | fault | ERROR | — | expected temperature cloud ACK 0x55 is missing at deadline. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-02 | 5 | 0x05 | SC | 2 | STM | fault | WARN | — | 0x55 arrives after transaction timeout/cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-03 | 5 | 0x05 | SC | 3 | STM | fault | ERROR | — | 0x55 status is upload failure (0). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-04 | 5 | 0x05 | SC | 4 | STM | fault | ERROR | — | 0x55 payload/message-ID length is malformed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-05 | 5 | 0x05 | SC | 5 | STM | fault | ERROR | — | ACK message ID does not match pending transaction where correlation exists. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-06 | 5 | 0x05 | SC | 6 | STM | fault | WARN | — | cloud warning 0x58 reports UNSTABLE; clears on OK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-07 | 5 | 0x05 | SC | 7 | STM | fault | ERROR | — | cloud warning 0x58 reports DOWN; clears on UNSTABLE/OK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-08 | 5 | 0x05 | SC | 8 | STM | fault | ERROR | — | 0x58 payload, quality, or context is invalid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-09 | 5 | 0x05 | SC | 9 | STM | fault | ERROR | — | WiFi status response has wrong payload length. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-10 | 5 | 0x05 | SC | 10 | STM | fault | WARN | — | WiFi status 0x00, disconnected; clears on connected. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-11 | 5 | 0x05 | SC | 11 | STM | fault | WARN | — | WiFi remains connecting/reconnecting beyond timeout after CMD_WIFI_RESP 0x13 or CMD_WIFI_CONNECTING 0x12. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-12 | 5 | 0x05 | SC | 12 | STM | fault | ERROR | — | WiFi status 0x03, explicit error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-13 | 5 | 0x05 | SC | 13 | STM | fault | ERROR | — | CMD_WIFI_RESP payload 0x15, suitable/configured AP not found. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-14 | 5 | 0x05 | SC | 14 | STM | fault | WARN | — | RSSI response timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-15 | 5 | 0x05 | SC | 15 | STM | fault | ERROR | — | RSSI response violates length/range contract. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-16 | 5 | 0x05 | SC | 16 | STM | fault | WARN | — | configured number of consecutive RSSI polls are missed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-19 | 5 | 0x05 | SC | 19 | STM | fault | ERROR | — | CMD_WIFI_RESP carries an undefined status payload. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-20 | 5 | 0x05 | SC | 20 | STM | fault | ERROR | — | standalone CMD_WIFI_FAILED (0x14) received; clears on later connected/connecting state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-01 | 6 | 0x06 | SS | 1 | STM | fault | CRITICAL | — | HardFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-02 | 6 | 0x06 | SS | 2 | STM | fault | CRITICAL | — | MemManage fault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-03 | 6 | 0x06 | SS | 3 | STM | fault | CRITICAL | — | BusFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-04 | 6 | 0x06 | SS | 4 | STM | fault | CRITICAL | — | UsageFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-05 | 6 | 0x06 | SS | 5 | STM | fault | CRITICAL | — | boot flags identify IWDG reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-06 | 6 | 0x06 | SS | 6 | STM | fault | CRITICAL | — | boot flags identify WWDG reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-07 | 6 | 0x06 | SS | 7 | STM | fault | CRITICAL | — | FreeRTOS stack-overflow hook fires. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-08 | 6 | 0x06 | SS | 8 | STM | fault | CRITICAL | — | FreeRTOS malloc-failed hook fires. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-09 | 6 | 0x06 | SS | 9 | STM | fault | CRITICAL | — | boot flags identify brownout reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-10 | 6 | 0x06 | SS | 10 | STM | fault | CRITICAL | — | configASSERT fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-11 | 6 | 0x06 | SS | 11 | STM | fault | CRITICAL | — | fatal application Error_Handler is entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-12 | 6 | 0x06 | SS | 12 | STM | fault | ERROR | — | diagnostic event ring overflows. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-01 | 7 | 0x07 | SO | 1 | STM | fault | ERROR | — | OTA available 0x50 payload malformed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-02 | 7 | 0x07 | SO | 2 | STM | fault | ERROR | — | OTA 0x52 phase is unknown. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-03 | 7 | 0x07 | SO | 3 | STM | fault | ERROR | — | STM cannot transmit OTA start 0x51/0x54. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-04 | 7 | 0x07 | SO | 4 | STM | fault | ERROR | — | expected OTA STARTING phase times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-05 | 7 | 0x07 | SO | 5 | STM | fault | ERROR | — | reserved for future OTA chunk-transfer silence timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-06 | 7 | 0x07 | SO | 6 | STM | fault | ERROR | — | OTA success 0x53 times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-07 | 7 | 0x07 | SO | 7 | STM | fault | ERROR | — | version/job-ID length fields violate bounds. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-08 | 7 | 0x07 | SO | 8 | STM | fault | WARN | — | OTA phase/success arrives in incompatible local state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-01 | 8 | 0x08 | PW | 1 | STM | fault | ERROR | — | MP2722 does not ACK at I2C4 0x3F. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-02 | 8 | 0x08 | PW | 2 | STM | fault | ERROR | — | MP2722 register read/write fails after detection. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-03 | 8 | 0x08 | PW | 3 | STM | fault | CRITICAL | — | REG12 charger watchdog fault. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-04 | 8 | 0x08 | PW | 4 | STM | fault | WARN | — | REG12 active die thermal regulation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-05 | 8 | 0x08 | PW | 5 | STM | fault | CRITICAL | — | REG13 charge input overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-06 | 8 | 0x08 | PW | 6 | STM | fault | ERROR | — | REG13 charge safety timer expired. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-07 | 8 | 0x08 | PW | 7 | STM | fault | CRITICAL | — | REG13 battery overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-08 | 8 | 0x08 | PW | 8 | STM | fault | CRITICAL | — | REG13 boost overload/short. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-09 | 8 | 0x08 | PW | 9 | STM | fault | ERROR | — | REG13 boost overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-10 | 8 | 0x08 | PW | 10 | STM | fault | CRITICAL | — | REG13 boost overtemperature. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-11 | 8 | 0x08 | PW | 11 | STM | fault | ERROR | — | REG13 battery too low for boost. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-12 | 8 | 0x08 | PW | 12 | STM | fault | CRITICAL | — | REG14 battery missing. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-13 | 8 | 0x08 | PW | 13 | STM | fault | ERROR | — | REG14 NTC missing. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-14 | 8 | 0x08 | PW | 14 | STM | fault | WARN | — | NTC1 WARM. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-15 | 8 | 0x08 | PW | 15 | STM | fault | WARN | — | NTC1 COOL. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-16 | 8 | 0x08 | PW | 16 | STM | fault | ERROR | — | NTC1 COLD. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-17 | 8 | 0x08 | PW | 17 | STM | fault | ERROR | — | NTC1 HOT. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-18 | 8 | 0x08 | PW | 18 | STM | fault | WARN | — | NTC2 WARM. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-19 | 8 | 0x08 | PW | 19 | STM | fault | WARN | — | NTC2 COOL. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-20 | 8 | 0x08 | PW | 20 | STM | fault | ERROR | — | NTC2 COLD. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-21 | 8 | 0x08 | PW | 21 | STM | fault | ERROR | — | NTC2 HOT. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-22 | 8 | 0x08 | PW | 22 | STM | fault | WARN | — | sustained input-current/input-voltage DPM throttling. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-23 | 8 | 0x08 | PW | 23 | STM | fault | ERROR | — | battery ADC4 DMA start/error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-24 | 8 | 0x08 | PW | 24 | STM | fault | CRITICAL | — | filtered battery voltage remains below validated critical threshold. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-25 | 8 | 0x08 | PW | 25 | STM | fault | ERROR | — | battery ADC is impossible, stale, or materially contradicts MP2722. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-01 | 9 | 0x09 | IB | 1 | STM | fault | ERROR | — | I2C1 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-02 | 9 | 0x09 | IB | 2 | STM | fault | ERROR | — | I2C1 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-03 | 9 | 0x09 | IB | 3 | STM | fault | ERROR | — | I2C1 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-04 | 9 | 0x09 | IB | 4 | STM | fault | ERROR | — | I2C1 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-11 | 9 | 0x09 | IB | 11 | STM | fault | ERROR | — | I2C3 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-12 | 9 | 0x09 | IB | 12 | STM | fault | ERROR | — | I2C3 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-13 | 9 | 0x09 | IB | 13 | STM | fault | ERROR | — | I2C3 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-14 | 9 | 0x09 | IB | 14 | STM | fault | ERROR | — | I2C3 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-21 | 9 | 0x09 | IB | 21 | STM | fault | ERROR | — | I2C4 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-22 | 9 | 0x09 | IB | 22 | STM | fault | ERROR | — | I2C4 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-23 | 9 | 0x09 | IB | 23 | STM | fault | ERROR | — | I2C4 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-24 | 9 | 0x09 | IB | 24 | STM | fault | ERROR | — | I2C4 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-01 | 10 | 0x0A | TM | 1 | STM | fault | WARN | — | no supported probe after PD4 power-on/settle during requested measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-02 | 10 | 0x0A | TM | 2 | STM | fault | ERROR | — | multiple mutually exclusive probe addresses respond. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-10 | 10 | 0x0A | TM | 10 | STM | fault | ERROR | — | TMP117 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-11 | 10 | 0x0A | TM | 11 | STM | fault | ERROR | — | TMP117 DEVICE_ID mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-12 | 10 | 0x0A | TM | 12 | STM | fault | ERROR | — | TMP117 register read fails after ACK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-13 | 10 | 0x0A | TM | 13 | STM | fault | ERROR | — | TMP117 measurement makes no progress by deadline. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-14 | 10 | 0x0A | TM | 14 | STM | fault | ERROR | — | TMP117 decoded value is NaN/impossible/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-20 | 10 | 0x0A | TM | 20 | STM | fault | ERROR | — | MLX90614 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-21 | 10 | 0x0A | TM | 21 | STM | fault | ERROR | — | MLX90614 paired ambient/object read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-22 | 10 | 0x0A | TM | 22 | STM | fault | ERROR | — | MLX90614 raw value is 0x0000/0xFFFF. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-23 | 10 | 0x0A | TM | 23 | STM | fault | WARN | — | MLX90614 acquisition has fewer than 8/8 successful sample pairs. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-24 | 10 | 0x0A | TM | 24 | STM | fault | ERROR | — | MLX90614 calculated value is NaN/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-30 | 10 | 0x0A | TM | 30 | STM | fault | ERROR | — | MLX90632 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-31 | 10 | 0x0A | TM | 31 | STM | fault | ERROR | — | MLX90632 register/RAM read failure. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-32 | 10 | 0x0A | TM | 32 | STM | fault | ERROR | — | MLX90632 DATA_RDY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-33 | 10 | 0x0A | TM | 33 | STM | fault | ERROR | — | MLX90632 brownout-reset status. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-34 | 10 | 0x0A | TM | 34 | STM | fault | ERROR | — | MLX90632 BUSY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-35 | 10 | 0x0A | TM | 35 | STM | fault | ERROR | — | MLX90632 EE_BUSY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-36 | 10 | 0x0A | TM | 36 | STM | fault | ERROR | — | MLX90632 EEPROM version/calibration invalid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-37 | 10 | 0x0A | TM | 37 | STM | fault | ERROR | — | MLX90632 calculated value is NaN/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-01 | 11 | 0x0B | AH | 1 | STM | fault | ERROR | — | TLV320DAC32 NACK while PD2 is commanded on. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-02 | 11 | 0x0B | AH | 2 | STM | fault | ERROR | — | required DAC initialization register write fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-03 | 11 | 0x0B | AH | 3 | STM | fault | ERROR | — | supported DAC critical-register readback mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-04 | 11 | 0x0B | AH | 4 | STM | fault | ERROR | — | SAI1 start/DMA operation fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-05 | 11 | 0x0B | AH | 5 | STM | fault | ERROR | — | SAI1 underrun/overrun/error callback. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-10 | 11 | 0x0B | AH | 10 | STM | fault | ERROR | — | MDF acquisition overflow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-11 | 11 | 0x0B | AH | 11 | STM | fault | ERROR | — | MDF resampling-filter overrun. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-12 | 11 | 0x0B | AH | 12 | STM | fault | ERROR | — | MDF clock absence. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-13 | 11 | 0x0B | AH | 13 | STM | fault | WARN | — | sustained MDF saturation/out-of-limit. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-14 | 11 | 0x0B | AH | 14 | STM | fault | CRITICAL | — | MDF short-circuit detection. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-15 | 11 | 0x0B | AH | 15 | STM | fault | ERROR | — | MDF DMA error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-16 | 11 | 0x0B | AH | 16 | STM | fault | ERROR | — | one microphone has sustained near-zero energy while the other is valid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-17 | 11 | 0x0B | AH | 17 | STM | fault | WARN | — | sustained microphone imbalance/clipping. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-01 | 12 | 0x0C | DP | 1 | STM | fault | CRITICAL | — | SPI1/display initialization fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-02 | 12 | 0x0C | DP | 2 | STM | fault | ERROR | — | SPI1 display transfer returns HAL error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-03 | 12 | 0x0C | DP | 3 | STM | fault | ERROR | — | SPI1 display transfer timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-04 | 12 | 0x0C | DP | 4 | STM | fault | ERROR | — | PD14 unexpectedly off while UI requires display. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-05 | 12 | 0x0C | DP | 5 | STM | fault | WARN | — | backlight PWM unexpectedly zero while UI requires visible output. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-06 | 12 | 0x0C | DP | 6 | STM | fault | ERROR | — | reset/init sequence fails to complete. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-01 | 13 | 0x0D | FM | 1 | STM | fault | CRITICAL | — | BSP_MEM_Init(0) fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-02 | 13 | 0x0D | FM | 2 | STM | fault | CRITICAL | — | JEDEC ID read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-03 | 13 | 0x0D | FM | 3 | STM | fault | CRITICAL | — | JEDEC ID mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-04 | 13 | 0x0D | FM | 4 | STM | fault | ERROR | — | flash status-register read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-05 | 13 | 0x0D | FM | 5 | STM | fault | CRITICAL | — | WIP remains set past operation timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-06 | 13 | 0x0D | FM | 6 | STM | fault | ERROR | — | WEL does not set after write-enable. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-07 | 13 | 0x0D | FM | 7 | STM | fault | ERROR | — | erase/program operation fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-08 | 13 | 0x0D | FM | 8 | STM | fault | CRITICAL | — | readback/CRC verification mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-09 | 13 | 0x0D | FM | 9 | STM | fault | CRITICAL | — | OCTOSPI memory-mapped-mode entry/sanity read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-01 | 14 | 0x0E | UB | 1 | STM | fault | ERROR | — | USB host/class initialization fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-02 | 14 | 0x0E | UB | 2 | STM | fault | WARN | — | target disconnects during active SpO2 flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-03 | 14 | 0x0E | UB | 3 | STM | fault | ERROR | — | enumerated PID is not 0x505C. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-04 | 14 | 0x0E | UB | 4 | STM | fault | ERROR | — | CDC TX submission fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-05 | 14 | 0x0E | UB | 5 | STM | fault | ERROR | — | CDC TX completion timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-06 | 14 | 0x0E | UB | 6 | STM | fault | ERROR | — | CDC RX submission fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-07 | 14 | 0x0E | UB | 7 | STM | fault | ERROR | — | CDC RX completion timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-08 | 14 | 0x0E | UB | 8 | STM | fault | ERROR | — | received SpO2 packet malformed/short. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-09 | 14 | 0x0E | UB | 9 | STM | fault | ERROR | — | SpO2/HR/status violates approved contract. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-10 | 14 | 0x0E | UB | 10 | STM | fault | ERROR | — | USB HCD transport error during active session. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-01 | 15 | 0x0F | IO | 1 | STM | fault | ERROR | — | PD4 temperature power state disagrees with active flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-02 | 15 | 0x0F | IO | 2 | STM | fault | ERROR | — | PD2 DAC/microphone power state disagrees with audio flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-03 | 15 | 0x0F | IO | 3 | STM | fault | ERROR | — | PD7/PD3 ESP power sequence disagrees with ESP state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-04 | 15 | 0x0F | IO | 4 | STM | fault | ERROR | — | display power/backlight state disagrees with UI sleep/wake state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-05 | 15 | 0x0F | IO | 5 | STM | fault | ERROR | — | PB14 peripheral-enable state disagrees with required state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-10 | 15 | 0x0F | IO | 10 | STM | fault | WARN | — | navigation button remains asserted beyond validated stuck-key duration. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-11 | 15 | 0x0F | IO | 11 | STM | fault | ERROR | — | wake/select event occurs but expected state transition times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PR-01 | 16 | 0x10 | PR | 1 | ESP | fault | ERROR | — | Provisioning / setup fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit / UART path
PR-02 | 16 | 0x10 | PR | 2 | ESP | fault | ERROR | — | Provisioning / setup fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
PR-03 | 16 | 0x10 | PR | 3 | ESP | fault | ERROR | — | Provisioning / setup fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
PR-04 | 16 | 0x10 | PR | 4 | ESP | fault | ERROR | — | Unknown / unsupported UART command | ESP detect → flash bucket → HTTPS | N/A (fault) | app_main default CMD
PR-05 | 16 | 0x10 | PR | 5 | ESP | fault | ERROR | — | Provisioning fault 5 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-01 | 17 | 0x11 | UA | 1 | ESP | fault | ERROR | — | UART read / link error | ESP detect → flash bucket → HTTPS | N/A (fault) | UART driver errors
UA-02 | 17 | 0x11 | UA | 2 | ESP | fault | ERROR | — | UART fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-03 | 17 | 0x11 | UA | 3 | ESP | fault | ERROR | — | UART fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-04 | 17 | 0x11 | UA | 4 | ESP | fault | ERROR | — | UART driver re-init / recovery | ESP detect → flash bucket → HTTPS | N/A (fault) | Multiple UART errors
UA-05 | 17 | 0x11 | UA | 5 | ESP | fault | ERROR | — | STM UART silence watchdog | ESP detect → flash bucket → HTTPS | N/A (fault) | No valid STM frames for timeout
SP-01 | 18 | 0x12 | SP | 1 | ESP | fault | ERROR | — | SPI RX fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-02 | 18 | 0x12 | SP | 2 | ESP | fault | ERROR | — | SPI RX fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-03 | 18 | 0x12 | SP | 3 | ESP | fault | ERROR | — | SPI fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-04 | 18 | 0x12 | SP | 4 | ESP | fault | ERROR | — | Audio ring overflow | ESP detect → flash bucket → HTTPS | N/A (fault) | Packetizer / ring buffer
SP-05 | 18 | 0x12 | SP | 5 | ESP | fault | ERROR | — | Packetizer fault | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio packetizer
WF-01 | 19 | 0x13 | WF | 1 | ESP | fault | ERROR | — | WiFi disconnected (any reason) | ESP detect → flash bucket → HTTPS | N/A (fault) | WiFi event
WF-02 | 19 | 0x13 | WF | 2 | ESP | fault | ERROR | — | No AP found | ESP detect → flash bucket → HTTPS | N/A (fault) | Scan / connect fail
WF-03 | 19 | 0x13 | WF | 3 | ESP | fault | ERROR | — | Lost IP after being online | ESP detect → flash bucket → HTTPS | N/A (fault) | Got-IP then lost
WF-04 | 19 | 0x13 | WF | 4 | ESP | fault | ERROR | — | WiFi fault 4 (reserved/defined) | ESP detect → flash bucket → HTTPS | N/A (fault) | Reserved
TL-01 | 20 | 0x14 | TL | 1 | ESP | fault | ERROR | — | Internet / TLS probe fail | ESP detect → flash bucket → HTTPS | N/A (fault) | AWS path / reachability
TL-02 | 20 | 0x14 | TL | 2 | ESP | fault | ERROR | — | TLS / reachability fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
TL-03 | 20 | 0x14 | TL | 3 | ESP | fault | ERROR | — | TLS / reachability fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
TL-04 | 20 | 0x14 | TL | 4 | ESP | fault | ERROR | — | TLS / MQTT TLS related | ESP detect → flash bucket → HTTPS | N/A (fault) | Handshake / socket
MQ-01 | 21 | 0x15 | MQ | 1 | ESP | fault | ERROR | — | MQTT connect / session fault | ESP detect → flash bucket → HTTPS | N/A (fault) | MQTT connect path
MQ-02 | 21 | 0x15 | MQ | 2 | ESP | fault | ERROR | — | MQTT fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-03 | 21 | 0x15 | MQ | 3 | ESP | fault | ERROR | — | MQTT fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-04 | 21 | 0x15 | MQ | 4 | ESP | fault | ERROR | — | MQTT fault 4 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-05 | 21 | 0x15 | MQ | 5 | ESP | fault | ERROR | — | MQTT fault 5 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-06 | 21 | 0x15 | MQ | 6 | ESP | fault | ERROR | — | MQTT fault 6 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-23 | 21 | 0x15 | MQ | 23 | ESP | fault | ERROR | — | Stream / publish path fault | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio/temp publish path
ES-01 | 22 | 0x16 | ES | 1 | ESP | fault | ERROR | — | OTA malloc fail | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
ES-02 | 22 | 0x16 | ES | 2 | ESP | fault | ERROR | — | Task create fail / coredump path | ESP detect → flash bucket → HTTPS | N/A (fault) | System
ES-03 | 22 | 0x16 | ES | 3 | ESP | fault | CRITICAL | — | Boot after WDT reset | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-04 | 22 | 0x16 | ES | 4 | ESP | fault | CRITICAL | — | Boot after panic | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-05 | 22 | 0x16 | ES | 5 | ESP | fault | CRITICAL | — | Boot after brownout | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-06 | 22 | 0x16 | ES | 6 | ESP | fault | ERROR | — | Core-dump present | ESP detect → flash bucket → HTTPS | N/A (fault) | Coredump partition
ES-07 | 22 | 0x16 | ES | 7 | ESP | fault | ERROR | — | Core-dump upload fail | ESP detect → flash bucket → HTTPS | N/A (fault) | HTTPS coredump
ES-08 | 22 | 0x16 | ES | 8 | ESP | fault | ERROR | — | Core-dump partition issue | ESP detect → flash bucket → HTTPS | N/A (fault) | Partition
ES-09 | 22 | 0x16 | ES | 9 | ESP | fault | WARN | — | Free heap < 40 KB | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >48KB
ES-10 | 22 | 0x16 | ES | 10 | ESP | fault | WARN | — | Min heap < 32 KB | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >40KB
ES-11 | 22 | 0x16 | ES | 11 | ESP | fault | WARN | — | mqtt_process stack HWM < 512 B | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >1024B
OT-01 | 23 | 0x17 | OT | 1 | ESP | fault | ERROR | — | ESP OTA fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-02 | 23 | 0x17 | OT | 2 | ESP | fault | ERROR | — | ESP OTA fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-03 | 23 | 0x17 | OT | 3 | ESP | fault | ERROR | — | ESP OTA fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-04 | 23 | 0x17 | OT | 4 | ESP | fault | ERROR | — | STM OTA fault | ESP detect → flash bucket → HTTPS | N/A (fault) | STM OTA path
PM-01 | 24 | 0x18 | PM | 1 | ESP | metric | INFO | bytes | Free heap | Periodic ~10 s health | value_u32 = free heap bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-02 | 24 | 0x18 | PM | 2 | ESP | metric | INFO | bytes | Min free heap (watermark) | Periodic ~10 s | value_u32 = min free heap | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-03 | 24 | 0x18 | PM | 3 | ESP | metric | INFO | bytes | Largest free block | Periodic ~10 s | value_u32 = largest block | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-04 | 24 | 0x18 | PM | 4 | ESP | metric | INFO | bytes | mqtt_process stack HWM | Periodic ~10 s | value_u32 = stack HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-05 | 24 | 0x18 | PM | 5 | ESP | metric | INFO | dBm | WiFi RSSI (signed) | Periodic + WiFi events | Interpret value as signed int32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-06 | 24 | 0x18 | PM | 6 | ESP | metric | INFO | count | WiFi reconnect count | Periodic + reconnect event | Monotonic counter | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-07 | 24 | 0x18 | PM | 7 | ESP | metric | INFO | count | MQTT disconnect count | Periodic + MQTT down | Monotonic counter | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-08 | 24 | 0x18 | PM | 8 | ESP | metric | INFO | s | Uptime since diag/health start | Periodic ~10 s | Seconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-09 | 24 | 0x18 | PM | 9 | ESP | metric | INFO | 0/1 | WiFi connected | Periodic + WiFi up/down | 1=connected, 0=not | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-10 | 24 | 0x18 | PM | 10 | ESP | metric | INFO | 0/1 | WiFi has IP | Periodic + disconnect→0 | 1=has IP | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-11 | 24 | 0x18 | PM | 11 | ESP | metric | INFO | 0/1 | Internet / AWS path OK | Periodic | 1=reachable | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-12 | 24 | 0x18 | PM | 12 | ESP | metric | INFO | 0/1 | MQTT connected | Periodic + MQTT up/down | 1=connected | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-13 | 24 | 0x18 | PM | 13 | ESP | metric | INFO | 0/1 | WiFi reconnect in progress | Periodic | 1=reconnecting | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-14 | 24 | 0x18 | PM | 14 | ESP | metric | INFO | index | Saved WiFi network index (1-based) | Periodic + connect | 0=none; 1=first saved SSID… | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-15 | 24 | 0x18 | PM | 15 | ESP | metric | INFO | hash | SSID FNV-1a u32 | Periodic + connect | Helper fingerprint; name in PM-29…37 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-16 | 24 | 0x18 | PM | 16 | ESP | metric | INFO | ms | MQTT keepalive ping RTT | On PINGRESP + periodic last | Exact RTT milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-17 | 24 | 0x18 | PM | 17 | ESP | metric | INFO | count | Temp cloud-ACK queue depth | Periodic | Pending ACKs | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-18 | 24 | 0x18 | PM | 18 | ESP | metric | INFO | count | MQTT outbox pending | Periodic | Outgoing slots | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-19 | 24 | 0x18 | PM | 19 | ESP | metric | INFO | count | MQTT event queue length | Periodic | 0 if unknown | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-20 | 24 | 0x18 | PM | 20 | ESP | metric | INFO | bytes | Internal RAM free | Periodic | MALLOC_CAP_INTERNAL | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-21 | 24 | 0x18 | PM | 21 | ESP | metric | INFO | bytes | Task diag_mgr stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-22 | 24 | 0x18 | PM | 22 | ESP | metric | INFO | bytes | Task uart_rx stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-23 | 24 | 0x18 | PM | 23 | ESP | metric | INFO | bytes | Task mqtt_diag stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-24 | 24 | 0x18 | PM | 24 | ESP | metric | INFO | channel | WiFi channel | Periodic + connect | AP primary channel | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-25 | 24 | 0x18 | PM | 25 | ESP | metric | INFO | ms | Last DNS resolve duration | On DNS + periodic last | Milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-26 | 24 | 0x18 | PM | 26 | ESP | metric | INFO | ms | Last TLS / AWS-path TLS duration | On TLS end / AWS probe | Milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-27 | 24 | 0x18 | PM | 27 | ESP | metric | INFO | 0/1 | AWS path reachable | On AWS probe + periodic | 1=OK | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-28 | 24 | 0x18 | PM | 28 | ESP | metric | INFO | 0/1 | SNTP / time trusted | Periodic | 1=trusted UTC | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-29 | 24 | 0x18 | PM | 29 | ESP | ssid_meta | INFO | chars | WiFi SSID length (0…32) | Health + WiFi connect | Use with PM-30…37 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-30 | 24 | 0x18 | PM | 30 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [0..3] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-31 | 24 | 0x18 | PM | 31 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [4..7] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-32 | 24 | 0x18 | PM | 32 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [8..11] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-33 | 24 | 0x18 | PM | 33 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [12..15] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-34 | 24 | 0x18 | PM | 34 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [16..19] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-35 | 24 | 0x18 | PM | 35 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [20..23] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-36 | 24 | 0x18 | PM | 36 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [24..27] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-37 | 24 | 0x18 | PM | 37 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [28..31] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-38 | 24 | 0x18 | PM | 38 | ESP | ssid_chunk | INFO | bytes | Heap total | Periodic ~10 s | MALLOC_CAP_DEFAULT total | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-39 | 24 | 0x18 | PM | 39 | ESP | ssid_chunk | INFO | bytes | Heap used | Periodic ~10 s | total − free | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-40 | 24 | 0x18 | PM | 40 | ESP | ssid_chunk | INFO | percent | Heap used % | Periodic ~10 s | 0…100 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-41 | 24 | 0x18 | PM | 41 | ESP | ssid_chunk | INFO | bytes | aws_iot_demo stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-42 | 24 | 0x18 | PM | 42 | ESP | ssid_chunk | INFO | bytes | mqtt_stream stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-43 | 24 | 0x18 | PM | 43 | ESP | ssid_chunk | INFO | bytes | audio_packetizer stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-44 | 24 | 0x18 | PM | 44 | ESP | ssid_chunk | INFO | bytes | spi_audio_rx stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-45 | 24 | 0x18 | PM | 45 | ESP | ssid_chunk | INFO | bytes | mqtt_process_task stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-46 | 24 | 0x18 | PM | 46 | ESP | ssid_chunk | INFO | bytes | temp_ack_retry stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-47 | 24 | 0x18 | PM | 47 | ESP | ssid_chunk | INFO | bytes | audio_ctrl_pub stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-48 | 24 | 0x18 | PM | 48 | ESP | ssid_chunk | INFO | bytes | mqtt_health_task stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-49 | 24 | 0x18 | PM | 49 | ESP | ssid_chunk | INFO | centi-°C | ESP chip die temperature (eFuse-corrected) | Periodic ~10 s | signed; ÷100 → °C; TEMP_CALIB in-app | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
SD-01 | 25 | 0x19 | SD | 1 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [0..3] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-02 | 25 | 0x19 | SD | 2 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [4..7] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-03 | 25 | 0x19 | SD | 3 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [8..11] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-04 | 25 | 0x19 | SD | 4 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [12..15] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-05 | 25 | 0x19 | SD | 5 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [16..19] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-06 | 25 | 0x19 | SD | 6 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [20..23] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-07 | 25 | 0x19 | SD | 7 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [24..27] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-08 | 25 | 0x19 | SD | 8 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [28..31] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-09 | 25 | 0x19 | SD | 9 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [32..35] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-10 | 25 | 0x19 | SD | 10 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [36..39] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-11 | 25 | 0x19 | SD | 11 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [40..43] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-12 | 25 | 0x19 | SD | 12 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [44..47] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-13 | 25 | 0x19 | SD | 13 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [48..51] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-14 | 25 | 0x19 | SD | 14 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [52..55] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-15 | 25 | 0x19 | SD | 15 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [56..59] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-16 | 25 | 0x19 | SD | 16 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [60..63] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-17 | 25 | 0x19 | SD | 17 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [64..67] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-18 | 25 | 0x19 | SD | 18 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [68..71] (batt + mcu_temp_c lo) | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-19 | 25 | 0x19 | SD | 19 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [72..73] (mcu_temp_c hi + ok + pad) | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame

## 03_STM_Faults
STM FAULT CODES
Families ST…IO · Sent as UART CMD 0x59 · ESP ACKs 0x5A · Stored in flash buckets
Code | family_id | hex | PREFIX | code_id | Owner | Kind | Severity | Unit | Meaning (what it means) | When / path | Value decode | Notes
ST-01 | 1 | 0x01 | ST | 1 | STM | fault | ERROR | — | LEN is below 2 or above the protocol maximum; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-02 | 1 | 0x01 | ST | 2 | STM | fault | ERROR | — | END byte is not 0x55; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-03 | 1 | 0x01 | ST | 3 | STM | fault | ERROR | — | received XOR does not match calculated XOR; clears on next valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-04 | 1 | 0x01 | ST | 4 | STM | fault | CRITICAL | — | raw RX buffer would overflow; clears after parser reset and a valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-05 | 1 | 0x01 | ST | 5 | STM | fault | WARN | — | START arrives during a partial frame and forces resynchronization; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-06 | 1 | 0x01 | ST | 6 | STM | fault | ERROR | — | partial frame exceeds assembly timeout; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-07 | 1 | 0x01 | ST | 7 | STM | fault | WARN | — | valid frame contains unsupported command ID; clears on supported command. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-08 | 1 | 0x01 | ST | 8 | STM | fault | ERROR | — | payload length violates command contract; clears on valid instance of that command. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-09 | 1 | 0x01 | ST | 9 | STM | fault | ERROR | — | nested length exceeds payload bounds; clears on valid command payload. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-10 | 1 | 0x01 | ST | 10 | STM | fault | ERROR | — | STM tries to build an oversized TX payload; clears after valid TX build. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
ST-11 | 1 | 0x01 | ST | 11 | STM | fault | CRITICAL | — | TX checksum self-verification fails; clears after verified build. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-01 | 2 | 0x02 | SU | 1 | STM | fault | ERROR | — | USART3 overrun (ORE); clears after flag recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-02 | 2 | 0x02 | SU | 2 | STM | fault | ERROR | — | framing error (FE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-03 | 2 | 0x02 | SU | 3 | STM | fault | WARN | — | noise error (NE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-04 | 2 | 0x02 | SU | 4 | STM | fault | ERROR | — | parity error (PE); clears after recovery and valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-05 | 2 | 0x02 | SU | 5 | STM | fault | WARN | — | receiver timeout/IDLE with incomplete frame; clears on valid frame. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-06 | 2 | 0x02 | SU | 6 | STM | fault | ERROR | — | USART3 RX DMA error if DMA is enabled later; clears after recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-07 | 2 | 0x02 | SU | 7 | STM | fault | ERROR | — | USART3 TX DMA error if DMA is enabled later; clears after successful TX. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-08 | 2 | 0x02 | SU | 8 | STM | fault | ERROR | — | finite-timeout packet TX returns HAL error/timeout; clears after successful TX. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-09 | 2 | 0x02 | SU | 9 | STM | fault | ERROR | — | ESP is silent for a generic request expecting a reply; clears when valid ESP traffic resumes. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SU-10 | 2 | 0x02 | SU | 10 | STM | fault | ERROR | — | validated RX queue is full and packet would be lost; clears after queue recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-01 | 3 | 0x03 | SA | 1 | STM | fault | ERROR | — | SPI2 audio transfer returns non-timeout HAL error; clears after successful block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-02 | 3 | 0x03 | SA | 2 | STM | fault | ERROR | — | SPI2 audio transfer times out; clears after successful block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-03 | 3 | 0x03 | SA | 3 | STM | fault | ERROR | — | SPI2 audio DMA error if enabled; clears after recovery. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-04 | 3 | 0x03 | SA | 4 | STM | fault | WARN | — | PCM transport queue is full and block is dropped; clears after sustained drain. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SA-05 | 3 | 0x03 | SA | 5 | STM | fault | ERROR | — | PCM block is null, zero, oversized, or misaligned; clears on valid block. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-01 | 4 | 0x04 | SF | 1 | STM | fault | ERROR | — | illegal stream state transition; clears after returning to IDLE/STOPPED. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-02 | 4 | 0x04 | SF | 2 | STM | fault | ERROR | — | START response timeout; clears on correlated success or cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-03 | 4 | 0x04 | SF | 3 | STM | fault | ERROR | — | START response reports failure; clears on later success/cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-04 | 4 | 0x04 | SF | 4 | STM | fault | ERROR | — | START response pet/audio/state fields do not match request; clears on correlated response. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-05 | 4 | 0x04 | SF | 5 | STM | fault | ERROR | — | STOP response timeout; clears on correlated success or forced reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-06 | 4 | 0x04 | SF | 6 | STM | fault | ERROR | — | STOP response reports failure; clears on later success/reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-07 | 4 | 0x04 | SF | 7 | STM | fault | ERROR | — | STOP response does not correlate; clears on correlated response. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-08 | 4 | 0x04 | SF | 8 | STM | fault | WARN | — | STM and ESP stream states diverge; clears when both agree. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SF-09 | 4 | 0x04 | SF | 9 | STM | fault | WARN | — | screen/mode exits while handshake is pending; clears after settlement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-01 | 5 | 0x05 | SC | 1 | STM | fault | ERROR | — | expected temperature cloud ACK 0x55 is missing at deadline. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-02 | 5 | 0x05 | SC | 2 | STM | fault | WARN | — | 0x55 arrives after transaction timeout/cancellation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-03 | 5 | 0x05 | SC | 3 | STM | fault | ERROR | — | 0x55 status is upload failure (0). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-04 | 5 | 0x05 | SC | 4 | STM | fault | ERROR | — | 0x55 payload/message-ID length is malformed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-05 | 5 | 0x05 | SC | 5 | STM | fault | ERROR | — | ACK message ID does not match pending transaction where correlation exists. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-06 | 5 | 0x05 | SC | 6 | STM | fault | WARN | — | cloud warning 0x58 reports UNSTABLE; clears on OK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-07 | 5 | 0x05 | SC | 7 | STM | fault | ERROR | — | cloud warning 0x58 reports DOWN; clears on UNSTABLE/OK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-08 | 5 | 0x05 | SC | 8 | STM | fault | ERROR | — | 0x58 payload, quality, or context is invalid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-09 | 5 | 0x05 | SC | 9 | STM | fault | ERROR | — | WiFi status response has wrong payload length. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-10 | 5 | 0x05 | SC | 10 | STM | fault | WARN | — | WiFi status 0x00, disconnected; clears on connected. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-11 | 5 | 0x05 | SC | 11 | STM | fault | WARN | — | WiFi remains connecting/reconnecting beyond timeout after CMD_WIFI_RESP 0x13 or CMD_WIFI_CONNECTING 0x12. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-12 | 5 | 0x05 | SC | 12 | STM | fault | ERROR | — | WiFi status 0x03, explicit error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-13 | 5 | 0x05 | SC | 13 | STM | fault | ERROR | — | CMD_WIFI_RESP payload 0x15, suitable/configured AP not found. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-14 | 5 | 0x05 | SC | 14 | STM | fault | WARN | — | RSSI response timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-15 | 5 | 0x05 | SC | 15 | STM | fault | ERROR | — | RSSI response violates length/range contract. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-16 | 5 | 0x05 | SC | 16 | STM | fault | WARN | — | configured number of consecutive RSSI polls are missed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-19 | 5 | 0x05 | SC | 19 | STM | fault | ERROR | — | CMD_WIFI_RESP carries an undefined status payload. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SC-20 | 5 | 0x05 | SC | 20 | STM | fault | ERROR | — | standalone CMD_WIFI_FAILED (0x14) received; clears on later connected/connecting state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-01 | 6 | 0x06 | SS | 1 | STM | fault | CRITICAL | — | HardFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-02 | 6 | 0x06 | SS | 2 | STM | fault | CRITICAL | — | MemManage fault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-03 | 6 | 0x06 | SS | 3 | STM | fault | CRITICAL | — | BusFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-04 | 6 | 0x06 | SS | 4 | STM | fault | CRITICAL | — | UsageFault handler entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-05 | 6 | 0x06 | SS | 5 | STM | fault | CRITICAL | — | boot flags identify IWDG reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-06 | 6 | 0x06 | SS | 6 | STM | fault | CRITICAL | — | boot flags identify WWDG reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-07 | 6 | 0x06 | SS | 7 | STM | fault | CRITICAL | — | FreeRTOS stack-overflow hook fires. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-08 | 6 | 0x06 | SS | 8 | STM | fault | CRITICAL | — | FreeRTOS malloc-failed hook fires. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-09 | 6 | 0x06 | SS | 9 | STM | fault | CRITICAL | — | boot flags identify brownout reset. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-10 | 6 | 0x06 | SS | 10 | STM | fault | CRITICAL | — | configASSERT fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-11 | 6 | 0x06 | SS | 11 | STM | fault | CRITICAL | — | fatal application Error_Handler is entered. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SS-12 | 6 | 0x06 | SS | 12 | STM | fault | ERROR | — | diagnostic event ring overflows. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-01 | 7 | 0x07 | SO | 1 | STM | fault | ERROR | — | OTA available 0x50 payload malformed. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-02 | 7 | 0x07 | SO | 2 | STM | fault | ERROR | — | OTA 0x52 phase is unknown. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-03 | 7 | 0x07 | SO | 3 | STM | fault | ERROR | — | STM cannot transmit OTA start 0x51/0x54. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-04 | 7 | 0x07 | SO | 4 | STM | fault | ERROR | — | expected OTA STARTING phase times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-05 | 7 | 0x07 | SO | 5 | STM | fault | ERROR | — | reserved for future OTA chunk-transfer silence timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-06 | 7 | 0x07 | SO | 6 | STM | fault | ERROR | — | OTA success 0x53 times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-07 | 7 | 0x07 | SO | 7 | STM | fault | ERROR | — | version/job-ID length fields violate bounds. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
SO-08 | 7 | 0x07 | SO | 8 | STM | fault | WARN | — | OTA phase/success arrives in incompatible local state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-01 | 8 | 0x08 | PW | 1 | STM | fault | ERROR | — | MP2722 does not ACK at I2C4 0x3F. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-02 | 8 | 0x08 | PW | 2 | STM | fault | ERROR | — | MP2722 register read/write fails after detection. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-03 | 8 | 0x08 | PW | 3 | STM | fault | CRITICAL | — | REG12 charger watchdog fault. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-04 | 8 | 0x08 | PW | 4 | STM | fault | WARN | — | REG12 active die thermal regulation. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-05 | 8 | 0x08 | PW | 5 | STM | fault | CRITICAL | — | REG13 charge input overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-06 | 8 | 0x08 | PW | 6 | STM | fault | ERROR | — | REG13 charge safety timer expired. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-07 | 8 | 0x08 | PW | 7 | STM | fault | CRITICAL | — | REG13 battery overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-08 | 8 | 0x08 | PW | 8 | STM | fault | CRITICAL | — | REG13 boost overload/short. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-09 | 8 | 0x08 | PW | 9 | STM | fault | ERROR | — | REG13 boost overvoltage. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-10 | 8 | 0x08 | PW | 10 | STM | fault | CRITICAL | — | REG13 boost overtemperature. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-11 | 8 | 0x08 | PW | 11 | STM | fault | ERROR | — | REG13 battery too low for boost. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-12 | 8 | 0x08 | PW | 12 | STM | fault | CRITICAL | — | REG14 battery missing. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-13 | 8 | 0x08 | PW | 13 | STM | fault | ERROR | — | REG14 NTC missing. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-14 | 8 | 0x08 | PW | 14 | STM | fault | WARN | — | NTC1 WARM. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-15 | 8 | 0x08 | PW | 15 | STM | fault | WARN | — | NTC1 COOL. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-16 | 8 | 0x08 | PW | 16 | STM | fault | ERROR | — | NTC1 COLD. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-17 | 8 | 0x08 | PW | 17 | STM | fault | ERROR | — | NTC1 HOT. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-18 | 8 | 0x08 | PW | 18 | STM | fault | WARN | — | NTC2 WARM. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-19 | 8 | 0x08 | PW | 19 | STM | fault | WARN | — | NTC2 COOL. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-20 | 8 | 0x08 | PW | 20 | STM | fault | ERROR | — | NTC2 COLD. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-21 | 8 | 0x08 | PW | 21 | STM | fault | ERROR | — | NTC2 HOT. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-22 | 8 | 0x08 | PW | 22 | STM | fault | WARN | — | sustained input-current/input-voltage DPM throttling. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-23 | 8 | 0x08 | PW | 23 | STM | fault | ERROR | — | battery ADC4 DMA start/error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-24 | 8 | 0x08 | PW | 24 | STM | fault | CRITICAL | — | filtered battery voltage remains below validated critical threshold. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
PW-25 | 8 | 0x08 | PW | 25 | STM | fault | ERROR | — | battery ADC is impossible, stale, or materially contradicts MP2722. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-01 | 9 | 0x09 | IB | 1 | STM | fault | ERROR | — | I2C1 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-02 | 9 | 0x09 | IB | 2 | STM | fault | ERROR | — | I2C1 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-03 | 9 | 0x09 | IB | 3 | STM | fault | ERROR | — | I2C1 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-04 | 9 | 0x09 | IB | 4 | STM | fault | ERROR | — | I2C1 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-11 | 9 | 0x09 | IB | 11 | STM | fault | ERROR | — | I2C3 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-12 | 9 | 0x09 | IB | 12 | STM | fault | ERROR | — | I2C3 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-13 | 9 | 0x09 | IB | 13 | STM | fault | ERROR | — | I2C3 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-14 | 9 | 0x09 | IB | 14 | STM | fault | ERROR | — | I2C3 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-21 | 9 | 0x09 | IB | 21 | STM | fault | ERROR | — | I2C4 bus error (BERR). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-22 | 9 | 0x09 | IB | 22 | STM | fault | ERROR | — | I2C4 arbitration lost (ARLO). | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-23 | 9 | 0x09 | IB | 23 | STM | fault | ERROR | — | I2C4 timeout / busy-stuck. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IB-24 | 9 | 0x09 | IB | 24 | STM | fault | ERROR | — | I2C4 DMA / internal error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-01 | 10 | 0x0A | TM | 1 | STM | fault | WARN | — | no supported probe after PD4 power-on/settle during requested measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-02 | 10 | 0x0A | TM | 2 | STM | fault | ERROR | — | multiple mutually exclusive probe addresses respond. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-10 | 10 | 0x0A | TM | 10 | STM | fault | ERROR | — | TMP117 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-11 | 10 | 0x0A | TM | 11 | STM | fault | ERROR | — | TMP117 DEVICE_ID mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-12 | 10 | 0x0A | TM | 12 | STM | fault | ERROR | — | TMP117 register read fails after ACK. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-13 | 10 | 0x0A | TM | 13 | STM | fault | ERROR | — | TMP117 measurement makes no progress by deadline. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-14 | 10 | 0x0A | TM | 14 | STM | fault | ERROR | — | TMP117 decoded value is NaN/impossible/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-20 | 10 | 0x0A | TM | 20 | STM | fault | ERROR | — | MLX90614 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-21 | 10 | 0x0A | TM | 21 | STM | fault | ERROR | — | MLX90614 paired ambient/object read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-22 | 10 | 0x0A | TM | 22 | STM | fault | ERROR | — | MLX90614 raw value is 0x0000/0xFFFF. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-23 | 10 | 0x0A | TM | 23 | STM | fault | WARN | — | MLX90614 acquisition has fewer than 8/8 successful sample pairs. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-24 | 10 | 0x0A | TM | 24 | STM | fault | ERROR | — | MLX90614 calculated value is NaN/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-30 | 10 | 0x0A | TM | 30 | STM | fault | ERROR | — | MLX90632 disappears during active measurement. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-31 | 10 | 0x0A | TM | 31 | STM | fault | ERROR | — | MLX90632 register/RAM read failure. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-32 | 10 | 0x0A | TM | 32 | STM | fault | ERROR | — | MLX90632 DATA_RDY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-33 | 10 | 0x0A | TM | 33 | STM | fault | ERROR | — | MLX90632 brownout-reset status. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-34 | 10 | 0x0A | TM | 34 | STM | fault | ERROR | — | MLX90632 BUSY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-35 | 10 | 0x0A | TM | 35 | STM | fault | ERROR | — | MLX90632 EE_BUSY timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-36 | 10 | 0x0A | TM | 36 | STM | fault | ERROR | — | MLX90632 EEPROM version/calibration invalid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
TM-37 | 10 | 0x0A | TM | 37 | STM | fault | ERROR | — | MLX90632 calculated value is NaN/outside approved range. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-01 | 11 | 0x0B | AH | 1 | STM | fault | ERROR | — | TLV320DAC32 NACK while PD2 is commanded on. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-02 | 11 | 0x0B | AH | 2 | STM | fault | ERROR | — | required DAC initialization register write fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-03 | 11 | 0x0B | AH | 3 | STM | fault | ERROR | — | supported DAC critical-register readback mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-04 | 11 | 0x0B | AH | 4 | STM | fault | ERROR | — | SAI1 start/DMA operation fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-05 | 11 | 0x0B | AH | 5 | STM | fault | ERROR | — | SAI1 underrun/overrun/error callback. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-10 | 11 | 0x0B | AH | 10 | STM | fault | ERROR | — | MDF acquisition overflow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-11 | 11 | 0x0B | AH | 11 | STM | fault | ERROR | — | MDF resampling-filter overrun. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-12 | 11 | 0x0B | AH | 12 | STM | fault | ERROR | — | MDF clock absence. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-13 | 11 | 0x0B | AH | 13 | STM | fault | WARN | — | sustained MDF saturation/out-of-limit. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-14 | 11 | 0x0B | AH | 14 | STM | fault | CRITICAL | — | MDF short-circuit detection. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-15 | 11 | 0x0B | AH | 15 | STM | fault | ERROR | — | MDF DMA error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-16 | 11 | 0x0B | AH | 16 | STM | fault | ERROR | — | one microphone has sustained near-zero energy while the other is valid. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
AH-17 | 11 | 0x0B | AH | 17 | STM | fault | WARN | — | sustained microphone imbalance/clipping. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-01 | 12 | 0x0C | DP | 1 | STM | fault | CRITICAL | — | SPI1/display initialization fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-02 | 12 | 0x0C | DP | 2 | STM | fault | ERROR | — | SPI1 display transfer returns HAL error. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-03 | 12 | 0x0C | DP | 3 | STM | fault | ERROR | — | SPI1 display transfer timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-04 | 12 | 0x0C | DP | 4 | STM | fault | ERROR | — | PD14 unexpectedly off while UI requires display. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-05 | 12 | 0x0C | DP | 5 | STM | fault | WARN | — | backlight PWM unexpectedly zero while UI requires visible output. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
DP-06 | 12 | 0x0C | DP | 6 | STM | fault | ERROR | — | reset/init sequence fails to complete. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-01 | 13 | 0x0D | FM | 1 | STM | fault | CRITICAL | — | BSP_MEM_Init(0) fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-02 | 13 | 0x0D | FM | 2 | STM | fault | CRITICAL | — | JEDEC ID read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-03 | 13 | 0x0D | FM | 3 | STM | fault | CRITICAL | — | JEDEC ID mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-04 | 13 | 0x0D | FM | 4 | STM | fault | ERROR | — | flash status-register read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-05 | 13 | 0x0D | FM | 5 | STM | fault | CRITICAL | — | WIP remains set past operation timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-06 | 13 | 0x0D | FM | 6 | STM | fault | ERROR | — | WEL does not set after write-enable. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-07 | 13 | 0x0D | FM | 7 | STM | fault | ERROR | — | erase/program operation fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-08 | 13 | 0x0D | FM | 8 | STM | fault | CRITICAL | — | readback/CRC verification mismatch. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
FM-09 | 13 | 0x0D | FM | 9 | STM | fault | CRITICAL | — | OCTOSPI memory-mapped-mode entry/sanity read fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-01 | 14 | 0x0E | UB | 1 | STM | fault | ERROR | — | USB host/class initialization fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-02 | 14 | 0x0E | UB | 2 | STM | fault | WARN | — | target disconnects during active SpO2 flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-03 | 14 | 0x0E | UB | 3 | STM | fault | ERROR | — | enumerated PID is not 0x505C. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-04 | 14 | 0x0E | UB | 4 | STM | fault | ERROR | — | CDC TX submission fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-05 | 14 | 0x0E | UB | 5 | STM | fault | ERROR | — | CDC TX completion timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-06 | 14 | 0x0E | UB | 6 | STM | fault | ERROR | — | CDC RX submission fails. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-07 | 14 | 0x0E | UB | 7 | STM | fault | ERROR | — | CDC RX completion timeout. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-08 | 14 | 0x0E | UB | 8 | STM | fault | ERROR | — | received SpO2 packet malformed/short. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-09 | 14 | 0x0E | UB | 9 | STM | fault | ERROR | — | SpO2/HR/status violates approved contract. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
UB-10 | 14 | 0x0E | UB | 10 | STM | fault | ERROR | — | USB HCD transport error during active session. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-01 | 15 | 0x0F | IO | 1 | STM | fault | ERROR | — | PD4 temperature power state disagrees with active flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-02 | 15 | 0x0F | IO | 2 | STM | fault | ERROR | — | PD2 DAC/microphone power state disagrees with audio flow. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-03 | 15 | 0x0F | IO | 3 | STM | fault | ERROR | — | PD7/PD3 ESP power sequence disagrees with ESP state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-04 | 15 | 0x0F | IO | 4 | STM | fault | ERROR | — | display power/backlight state disagrees with UI sleep/wake state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-05 | 15 | 0x0F | IO | 5 | STM | fault | ERROR | — | PB14 peripheral-enable state disagrees with required state. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-10 | 15 | 0x0F | IO | 10 | STM | fault | WARN | — | navigation button remains asserted beyond validated stuck-key duration. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan
IO-11 | 15 | 0x0F | IO | 11 | STM | fault | ERROR | — | wake/select event occurs but expected state transition times out. | STM detects → UART 0x59 → ESP flash → HTTPS | N/A (fault; session/seq are identity) | STM Error Book / COMPONENT plan

## 04_ESP_Faults
ESP FAULT CODES
Families PR…OT · diag_report → same 16-byte records → HTTPS buckets
Code | family_id | hex | PREFIX | code_id | Owner | Kind | Severity | Unit | Meaning (what it means) | When / path | Value decode | Notes
PR-01 | 16 | 0x10 | PR | 1 | ESP | fault | ERROR | — | Provisioning / setup fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit / UART path
PR-02 | 16 | 0x10 | PR | 2 | ESP | fault | ERROR | — | Provisioning / setup fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
PR-03 | 16 | 0x10 | PR | 3 | ESP | fault | ERROR | — | Provisioning / setup fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
PR-04 | 16 | 0x10 | PR | 4 | ESP | fault | ERROR | — | Unknown / unsupported UART command | ESP detect → flash bucket → HTTPS | N/A (fault) | app_main default CMD
PR-05 | 16 | 0x10 | PR | 5 | ESP | fault | ERROR | — | Provisioning fault 5 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-01 | 17 | 0x11 | UA | 1 | ESP | fault | ERROR | — | UART read / link error | ESP detect → flash bucket → HTTPS | N/A (fault) | UART driver errors
UA-02 | 17 | 0x11 | UA | 2 | ESP | fault | ERROR | — | UART fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-03 | 17 | 0x11 | UA | 3 | ESP | fault | ERROR | — | UART fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
UA-04 | 17 | 0x11 | UA | 4 | ESP | fault | ERROR | — | UART driver re-init / recovery | ESP detect → flash bucket → HTTPS | N/A (fault) | Multiple UART errors
UA-05 | 17 | 0x11 | UA | 5 | ESP | fault | ERROR | — | STM UART silence watchdog | ESP detect → flash bucket → HTTPS | N/A (fault) | No valid STM frames for timeout
SP-01 | 18 | 0x12 | SP | 1 | ESP | fault | ERROR | — | SPI RX fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-02 | 18 | 0x12 | SP | 2 | ESP | fault | ERROR | — | SPI RX fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-03 | 18 | 0x12 | SP | 3 | ESP | fault | ERROR | — | SPI fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio SPI path
SP-04 | 18 | 0x12 | SP | 4 | ESP | fault | ERROR | — | Audio ring overflow | ESP detect → flash bucket → HTTPS | N/A (fault) | Packetizer / ring buffer
SP-05 | 18 | 0x12 | SP | 5 | ESP | fault | ERROR | — | Packetizer fault | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio packetizer
WF-01 | 19 | 0x13 | WF | 1 | ESP | fault | ERROR | — | WiFi disconnected (any reason) | ESP detect → flash bucket → HTTPS | N/A (fault) | WiFi event
WF-02 | 19 | 0x13 | WF | 2 | ESP | fault | ERROR | — | No AP found | ESP detect → flash bucket → HTTPS | N/A (fault) | Scan / connect fail
WF-03 | 19 | 0x13 | WF | 3 | ESP | fault | ERROR | — | Lost IP after being online | ESP detect → flash bucket → HTTPS | N/A (fault) | Got-IP then lost
WF-04 | 19 | 0x13 | WF | 4 | ESP | fault | ERROR | — | WiFi fault 4 (reserved/defined) | ESP detect → flash bucket → HTTPS | N/A (fault) | Reserved
TL-01 | 20 | 0x14 | TL | 1 | ESP | fault | ERROR | — | Internet / TLS probe fail | ESP detect → flash bucket → HTTPS | N/A (fault) | AWS path / reachability
TL-02 | 20 | 0x14 | TL | 2 | ESP | fault | ERROR | — | TLS / reachability fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
TL-03 | 20 | 0x14 | TL | 3 | ESP | fault | ERROR | — | TLS / reachability fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
TL-04 | 20 | 0x14 | TL | 4 | ESP | fault | ERROR | — | TLS / MQTT TLS related | ESP detect → flash bucket → HTTPS | N/A (fault) | Handshake / socket
MQ-01 | 21 | 0x15 | MQ | 1 | ESP | fault | ERROR | — | MQTT connect / session fault | ESP detect → flash bucket → HTTPS | N/A (fault) | MQTT connect path
MQ-02 | 21 | 0x15 | MQ | 2 | ESP | fault | ERROR | — | MQTT fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-03 | 21 | 0x15 | MQ | 3 | ESP | fault | ERROR | — | MQTT fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-04 | 21 | 0x15 | MQ | 4 | ESP | fault | ERROR | — | MQTT fault 4 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-05 | 21 | 0x15 | MQ | 5 | ESP | fault | ERROR | — | MQTT fault 5 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-06 | 21 | 0x15 | MQ | 6 | ESP | fault | ERROR | — | MQTT fault 6 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP emit
MQ-23 | 21 | 0x15 | MQ | 23 | ESP | fault | ERROR | — | Stream / publish path fault | ESP detect → flash bucket → HTTPS | N/A (fault) | Audio/temp publish path
ES-01 | 22 | 0x16 | ES | 1 | ESP | fault | ERROR | — | OTA malloc fail | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
ES-02 | 22 | 0x16 | ES | 2 | ESP | fault | ERROR | — | Task create fail / coredump path | ESP detect → flash bucket → HTTPS | N/A (fault) | System
ES-03 | 22 | 0x16 | ES | 3 | ESP | fault | CRITICAL | — | Boot after WDT reset | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-04 | 22 | 0x16 | ES | 4 | ESP | fault | CRITICAL | — | Boot after panic | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-05 | 22 | 0x16 | ES | 5 | ESP | fault | CRITICAL | — | Boot after brownout | ESP detect → flash bucket → HTTPS | N/A (fault) | Boot reason
ES-06 | 22 | 0x16 | ES | 6 | ESP | fault | ERROR | — | Core-dump present | ESP detect → flash bucket → HTTPS | N/A (fault) | Coredump partition
ES-07 | 22 | 0x16 | ES | 7 | ESP | fault | ERROR | — | Core-dump upload fail | ESP detect → flash bucket → HTTPS | N/A (fault) | HTTPS coredump
ES-08 | 22 | 0x16 | ES | 8 | ESP | fault | ERROR | — | Core-dump partition issue | ESP detect → flash bucket → HTTPS | N/A (fault) | Partition
ES-09 | 22 | 0x16 | ES | 9 | ESP | fault | WARN | — | Free heap < 40 KB | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >48KB
ES-10 | 22 | 0x16 | ES | 10 | ESP | fault | WARN | — | Min heap < 32 KB | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >40KB
ES-11 | 22 | 0x16 | ES | 11 | ESP | fault | WARN | — | mqtt_process stack HWM < 512 B | ESP detect → flash bucket → HTTPS | N/A (fault) | Threshold + hysteresis clear >1024B
OT-01 | 23 | 0x17 | OT | 1 | ESP | fault | ERROR | — | ESP OTA fault 1 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-02 | 23 | 0x17 | OT | 2 | ESP | fault | ERROR | — | ESP OTA fault 2 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-03 | 23 | 0x17 | OT | 3 | ESP | fault | ERROR | — | ESP OTA fault 3 | ESP detect → flash bucket → HTTPS | N/A (fault) | ESP OTA
OT-04 | 23 | 0x17 | OT | 4 | ESP | fault | ERROR | — | STM OTA fault | ESP detect → flash bucket → HTTPS | N/A (fault) | STM OTA path

## 05_PM_Metrics
ESP PM METRICS + SSID
Family PM (0x18) · Integers in value_u32 · SSID name = PM-29 + PM-30…37
Code | family_id | hex | PREFIX | code_id | Owner | Kind | Severity | Unit | Meaning (what it means) | When / path | Value decode | Notes
PM-01 | 24 | 0x18 | PM | 1 | ESP | metric | INFO | bytes | Free heap | Periodic ~10 s health | value_u32 = free heap bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-02 | 24 | 0x18 | PM | 2 | ESP | metric | INFO | bytes | Min free heap (watermark) | Periodic ~10 s | value_u32 = min free heap | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-03 | 24 | 0x18 | PM | 3 | ESP | metric | INFO | bytes | Largest free block | Periodic ~10 s | value_u32 = largest block | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-04 | 24 | 0x18 | PM | 4 | ESP | metric | INFO | bytes | mqtt_process stack HWM | Periodic ~10 s | value_u32 = stack HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-05 | 24 | 0x18 | PM | 5 | ESP | metric | INFO | dBm | WiFi RSSI (signed) | Periodic + WiFi events | Interpret value as signed int32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-06 | 24 | 0x18 | PM | 6 | ESP | metric | INFO | count | WiFi reconnect count | Periodic + reconnect event | Monotonic counter | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-07 | 24 | 0x18 | PM | 7 | ESP | metric | INFO | count | MQTT disconnect count | Periodic + MQTT down | Monotonic counter | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-08 | 24 | 0x18 | PM | 8 | ESP | metric | INFO | s | Uptime since diag/health start | Periodic ~10 s | Seconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-09 | 24 | 0x18 | PM | 9 | ESP | metric | INFO | 0/1 | WiFi connected | Periodic + WiFi up/down | 1=connected, 0=not | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-10 | 24 | 0x18 | PM | 10 | ESP | metric | INFO | 0/1 | WiFi has IP | Periodic + disconnect→0 | 1=has IP | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-11 | 24 | 0x18 | PM | 11 | ESP | metric | INFO | 0/1 | Internet / AWS path OK | Periodic | 1=reachable | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-12 | 24 | 0x18 | PM | 12 | ESP | metric | INFO | 0/1 | MQTT connected | Periodic + MQTT up/down | 1=connected | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-13 | 24 | 0x18 | PM | 13 | ESP | metric | INFO | 0/1 | WiFi reconnect in progress | Periodic | 1=reconnecting | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-14 | 24 | 0x18 | PM | 14 | ESP | metric | INFO | index | Saved WiFi network index (1-based) | Periodic + connect | 0=none; 1=first saved SSID… | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-15 | 24 | 0x18 | PM | 15 | ESP | metric | INFO | hash | SSID FNV-1a u32 | Periodic + connect | Helper fingerprint; name in PM-29…37 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-16 | 24 | 0x18 | PM | 16 | ESP | metric | INFO | ms | MQTT keepalive ping RTT | On PINGRESP + periodic last | Exact RTT milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-17 | 24 | 0x18 | PM | 17 | ESP | metric | INFO | count | Temp cloud-ACK queue depth | Periodic | Pending ACKs | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-18 | 24 | 0x18 | PM | 18 | ESP | metric | INFO | count | MQTT outbox pending | Periodic | Outgoing slots | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-19 | 24 | 0x18 | PM | 19 | ESP | metric | INFO | count | MQTT event queue length | Periodic | 0 if unknown | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-20 | 24 | 0x18 | PM | 20 | ESP | metric | INFO | bytes | Internal RAM free | Periodic | MALLOC_CAP_INTERNAL | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-21 | 24 | 0x18 | PM | 21 | ESP | metric | INFO | bytes | Task diag_mgr stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-22 | 24 | 0x18 | PM | 22 | ESP | metric | INFO | bytes | Task uart_rx stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-23 | 24 | 0x18 | PM | 23 | ESP | metric | INFO | bytes | Task mqtt_diag stack HWM | Periodic | uxTaskGetStackHighWaterMark | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-24 | 24 | 0x18 | PM | 24 | ESP | metric | INFO | channel | WiFi channel | Periodic + connect | AP primary channel | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-25 | 24 | 0x18 | PM | 25 | ESP | metric | INFO | ms | Last DNS resolve duration | On DNS + periodic last | Milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-26 | 24 | 0x18 | PM | 26 | ESP | metric | INFO | ms | Last TLS / AWS-path TLS duration | On TLS end / AWS probe | Milliseconds | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-27 | 24 | 0x18 | PM | 27 | ESP | metric | INFO | 0/1 | AWS path reachable | On AWS probe + periodic | 1=OK | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-28 | 24 | 0x18 | PM | 28 | ESP | metric | INFO | 0/1 | SNTP / time trusted | Periodic | 1=trusted UTC | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-29 | 24 | 0x18 | PM | 29 | ESP | ssid_meta | INFO | chars | WiFi SSID length (0…32) | Health + WiFi connect | Use with PM-30…37 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-30 | 24 | 0x18 | PM | 30 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [0..3] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-31 | 24 | 0x18 | PM | 31 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [4..7] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-32 | 24 | 0x18 | PM | 32 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [8..11] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-33 | 24 | 0x18 | PM | 33 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [12..15] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-34 | 24 | 0x18 | PM | 34 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [16..19] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-35 | 24 | 0x18 | PM | 35 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [20..23] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-36 | 24 | 0x18 | PM | 36 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [24..27] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-37 | 24 | 0x18 | PM | 37 | ESP | ssid_chunk | INFO | bytes[4] | SSID ASCII bytes [28..31] | Health + WiFi connect | LE bytes in value_u32 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-38 | 24 | 0x18 | PM | 38 | ESP | ssid_chunk | INFO | bytes | Heap total | Periodic ~10 s | MALLOC_CAP_DEFAULT total | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-39 | 24 | 0x18 | PM | 39 | ESP | ssid_chunk | INFO | bytes | Heap used | Periodic ~10 s | total − free | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-40 | 24 | 0x18 | PM | 40 | ESP | ssid_chunk | INFO | percent | Heap used % | Periodic ~10 s | 0…100 | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-41 | 24 | 0x18 | PM | 41 | ESP | ssid_chunk | INFO | bytes | aws_iot_demo stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-42 | 24 | 0x18 | PM | 42 | ESP | ssid_chunk | INFO | bytes | mqtt_stream stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-43 | 24 | 0x18 | PM | 43 | ESP | ssid_chunk | INFO | bytes | audio_packetizer stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-44 | 24 | 0x18 | PM | 44 | ESP | ssid_chunk | INFO | bytes | spi_audio_rx stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-45 | 24 | 0x18 | PM | 45 | ESP | ssid_chunk | INFO | bytes | mqtt_process_task stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-46 | 24 | 0x18 | PM | 46 | ESP | ssid_chunk | INFO | bytes | temp_ack_retry stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-47 | 24 | 0x18 | PM | 47 | ESP | ssid_chunk | INFO | bytes | audio_ctrl_pub stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-48 | 24 | 0x18 | PM | 48 | ESP | ssid_chunk | INFO | bytes | mqtt_health_task stack HWM | Periodic ~10 s | task HWM bytes | flags ¦= 0x04 METRIC; value = session¦(seq<<16)
PM-49 | 24 | 0x18 | PM | 49 | ESP | ssid_chunk | INFO | centi-°C | ESP chip die temperature (eFuse-corrected) | Periodic ~10 s | signed; ÷100 → °C; TEMP_CALIB in-app | flags ¦= 0x04 METRIC; value = session¦(seq<<16)

## 06_SD_Status_Chunks
STM STATUS CHUNKS (0x5B → SD)
Family SD (0x19) · No ACK · 19 chunks rebuild 74-byte payload · See sheet 07 for field map
Code | family_id | hex | PREFIX | code_id | Owner | Kind | Severity | Unit | Meaning (what it means) | When / path | Value decode | Notes
SD-01 | 25 | 0x19 | SD | 1 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [0..3] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-02 | 25 | 0x19 | SD | 2 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [4..7] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-03 | 25 | 0x19 | SD | 3 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [8..11] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-04 | 25 | 0x19 | SD | 4 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [12..15] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-05 | 25 | 0x19 | SD | 5 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [16..19] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-06 | 25 | 0x19 | SD | 6 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [20..23] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-07 | 25 | 0x19 | SD | 7 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [24..27] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-08 | 25 | 0x19 | SD | 8 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [28..31] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-09 | 25 | 0x19 | SD | 9 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [32..35] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-10 | 25 | 0x19 | SD | 10 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [36..39] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-11 | 25 | 0x19 | SD | 11 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [40..43] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-12 | 25 | 0x19 | SD | 12 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [44..47] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-13 | 25 | 0x19 | SD | 13 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [48..51] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-14 | 25 | 0x19 | SD | 14 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [52..55] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-15 | 25 | 0x19 | SD | 15 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [56..59] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-16 | 25 | 0x19 | SD | 16 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [60..63] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-17 | 25 | 0x19 | SD | 17 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [64..67] | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-18 | 25 | 0x19 | SD | 18 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [68..71] (batt + mcu_temp_c lo) | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame
SD-19 | 25 | 0x19 | SD | 19 | ESP | status_chunk | INFO | bytes[4] | STM 0x5B V4 status payload bytes [72..73] (mcu_temp_c hi + ok + pad) | STM UART CMD 0x5B LEN=0x4C (no ACK) → ESP splits → flash | LE 4 bytes; concat SD-01…19 → take 74 bytes | Same event_time_s for all 19 chunks of one frame

## 07_Status_0x5B_Fields
STM 0x5B PAYLOAD FIELD MAP
After server reassembles SD-01…SD-16 → payload74[0:63] · All multi-byte fields little-endian
Offset | Type | Field name | Meaning / enum | Notes | Screen | Name | Button | Name | Action | Name
0 | u16 | session_id | STM boot session (like 0x59) | LE | 0 | LOGO | 255 | none | 0 | none
2 | u16 | seq | Status sequence (independent of alert seq) | LE | 1 | WIFI_Connect | 0 | select | 1 | screen_enter
4 | u8 | kind | 0=health, 1=UI |  | 2 | Home | 1 | up | 10 | Home select
5 | u8 | screen | 0 LOGO … 10 Charging (see enums on right) |  | 3 | WIFI_QR | 2 | right | 11 | Home back/home
6 | u8 | wifi_status | 0 Connecting, 1 Connected, 2 Error |  | 4 | TempProbe | 3 | down | 20 | Exam select
7 | u8 | wifi_rssi_cat | 1…5 bars (5=no signal) |  | 5 | TempProcess | 4 | left | 30 | Temp sample
8 | u8 | mqtt | 0/1 |  | 6 | TempResult | 5 | back | 31 | Temp save
9 | u8 | audio_streaming | 0/1 |  | 7 | ManualID | 6 | home | 40 | stream start
10 | u8 | audio_type | currentAudioTypeCode |  | 8 | Examination |  |  | 41 | stream stop
11 | u8 | button | 255=none; 0 select … 6 home |  | 9 | Auscultation |  |  | 42 | Aus back
12 | u8 | action | 0 none, 1 screen_enter, 10/11/20/30/31/40/41/42/50 |  | 10 | Charging |  |  | 50 | QR nav
13 | u8 | button_age_s | Seconds since last button/action (cap 255)
14 | u32 | heap_total_B | STM FreeRTOS heap total bytes | LE
18 | u32 | heap_free_B | Free bytes | LE
22 | u32 | heap_used_B | Used bytes | LE
26 | u8 | heap_pct | 0…100
27 | u32 | upload_kbps | kb/s integer | LE
31 | u8 | mqtt_q_pct | 0…100
32 | u8 | ring_pct | 0…100
33 | u16 | jitter_mad_ms | ms | LE
35 | u16 | esp32_used_B | Esp32Task stack used | LE
37 | u16 | esp32_total_B | Esp32Task stack total | LE
39 | u16 | gui_used_B | GUI task stack used | LE
41 | u16 | gui_total_B | GUI task stack total | LE
43 | u16 | audio_used_B | Audio task stack used | LE
45 | u16 | audio_total_B | Audio task stack total | LE
47 | u16 | usb_used_B | USB task stack used | LE
49 | u16 | usb_total_B | USB task stack total | LE
51 | u16 | temp1_used_B | Temp1 task stack used | LE
53 | u16 | temp1_total_B | Temp1 task stack total | LE
55 | u16 | vitals_used_B | Vitals task stack used | LE
57 | u16 | vitals_total_B | Vitals task stack total | LE
59 | u16 | mp2722_used_B | MP2722 task stack used | LE
61 | u16 | mp2722_total_B | MP2722 task stack total | LE

## 08_Flow_Constants
END-TO-END FLOW & CONSTANTS
Same for faults, PM metrics, SSID chunks, and SD status
Step | What happens
1 Detect | STM fault/status OR ESP fault/metric/health/WiFi/MQTT/ping/DNS/TLS
2 Encode | 16-byte record: family, code, repeat, flags, event_time_s, session, seq, CRC
3 Store | Append to active 32 KB bucket (0..3) inside 128 KB diag_log
4 Seal | When full → seal footer @ 0x7FF8; switch to next free bucket
5 Upload | HTTPS POST body=32768, X-Diag-Schema=bucket-v1, X-Bucket-Id=0..3
6 Erase | HTTP 2xx → erase that bucket → free for new logs
7 Busy | If all 4 sealed → new writes pause; STM 0x59 gets BUSY ACK
Constants for server
Key | Value
Bucket size / body length | 32768
format_ver | 3
X-Bucket-Id | 0 … 3
Seal footer offset | 0x7FF8
Record size | 16 bytes
Max records / bucket | 2045
METRIC flag | flags & 0x04
Metric value | session ¦ (sequence << 16)
SSID rebuild | len=PM-29; bytes=concat(PM-30…37); ssid=bytes[0:len]
0x5B rebuild | concat(SD-01…19) → first 74 bytes
STM alert CMD | 0x59 (ACK 0x5A)
STM status CMD | 0x5B (no ACK), LEN=0x41
Success HTTP | any 2xx (e.g. 204)
SSID example: HomeWiFi
Code | Value
PM-29 | 8
PM-30 | 0x656D6F48 → bytes H o m e
PM-31 | 0x69466957 → bytes W i F i
PM-32…37 | 0
Result | "HomeWiFi"